Table of Contents

# How the CPAP algorithms work — plain-English guide

This document explains, in ordinary language, what each of the important algorithms in this firmware actually does and why it is built that way. It is aimed at someone who needs to work on the code, tune it on the bench, or understand a strange trace, without first reading 3000 lines of C.

It is a companion to pressure_control_block_diagram.md, which has the signal-flow diagrams and the exact variable names. This document is the "why" and the "in words"; that one is the "what connects to what".

Written against the working tree on 2026-09-21. Default values quoted here are the compiled-in defaults; almost all of them are volatile globals you can change live from the debugger or over USB, so a running unit may differ.

## 1. The big picture

The device has one job: keep the air pressure at the patient's mask at the pressure the clinician set, while the patient breathes in and out of the same hose.

That is harder than it sounds, for three reasons:

- There is no sensor at the mask. The pressure sensor is inside the device. Air flowing down the hose loses pressure on the way, so what the device measures is not what the patient gets.

- The patient is a disturbance. Every breath in sucks air out of the circuit and pressure sags; every breath out pushes air back and pressure rises. The controller has to fight this several times a minute, forever.

- The blower is slow and non-linear. You cannot change the pressure directly. You can only change the fan speed, and the fan takes time to change speed, and the relationship between speed and pressure is a curve, not a straight line.

The control chain, end to end:

therapy setting (e.g. 8 cmH2O)
   -> therapy state machine (ramp, EPR, APAP, mask-off)
   -> pressure setpoint
   -> pressure PID + feed-forward        [CM7, 200 times a second]
   -> target motor RPM
   -> shared memory mailbox
   -> motor field-oriented control       [CM4]
   -> blower
   -> hose -> mask -> patient
   -> pressure and flow sensors -> back to the top

### Two processors

This is an STM32H747, which has two CPU cores in one chip, and the work is split between them.

- CM7 (the big core) runs everything you would call "the product": the screen, the sensors, the pressure control, therapy logic, safety, logging.

- CM4 (the small core) runs only the motor. It does field-oriented control (FOC) of the blower's brushless motor — a fast, dedicated job that must never be interrupted by a screen redraw.

They talk through a small block of shared memory (a "mailbox") at address 0x38000000, defined in Common/Inc/blower_ipc.h. CM7 writes commands into it (start, stop, set speed); CM4 writes status back (actual RPM, faults, bus voltage). Each side bumps a sequence counter so the other can tell a new message from a repeated one.

The important consequence: CM7 never controls the motor directly. It asks for an RPM, and CM4 delivers it. If the mailbox stops being serviced, CM7 notices through a heartbeat check and faults.

## 2. Who runs what, and how often

Everything on CM7 runs under FreeRTOS. The tasks, in priority order:

| Task | Rate | Priority | Job |
| --- | --- | --- | --- |
| SafetyTask | 200 Hz | Realtime | Watches pressure and the CM4 heartbeat. Can stop therapy. |
| SensorTask | 200 Hz | Above normal | Reads sensors, runs leak, therapy and the pressure PID. |
| GuiTask | ~60 Hz | Normal | TouchGFX screen drawing, rotary knob. |
| NfcTask | 100 Hz | Normal | NFC tag reading. |
| SystemTask | — | Below normal | USB logging, bench commands. |
| ClimateTask | 4 Hz | Below normal | Temperature/humidity sensors, humidifier loop. |
| SelfTestTask | — | Low | SD card and memory tests. |

SafetyTask deliberately sits above the control loop. If the control loop ever hangs or misbehaves, the thing that can shut it down still runs.

Inside SensorTask, one tick (5 ms) does this, in this exact order:

- Read the pressure sensors over SPI.

- Convert the flow sensor's differential pressure into a flow number.

- Update the leak estimate and patient flow.

- Therapy_Update() — state machine, EPR, APAP, mask-off. This sets the pressure setpoint.

- BlowerCtrl_Step() — the pressure PID. This sets the motor RPM.

- Push telemetry to USB and to the ESP32.

The order matters. Therapy runs before the PID so that the setpoint the PID sees is the one decided this tick, not last tick's.

## 3. What the sensors actually measure

Three AMS5935 pressure sensors sit on the SPI bus:

- PS — patient/mask pressure. An absolute-ish reading of the pressure in the circuit, measured inside the device. This is the feedback for the pressure loop. Reported in millibar, converted to cmH₂O (1 cmH₂O = 0.980665 mbar).

- FS — flow. This does not measure flow. It measures the tiny pressure difference across a restriction in the air path (±5 mbar full scale). Flow is worked out from that difference through a calibration table.

- AS — ambient. Room pressure, used for compensation.

Two practical facts that shape the code:

- SPI is shared, so PS is not refreshed on every single 5 ms tick. The code publishes a g_ps_sample_fresh flag, and anything that differentiates pressure (the D term, the dip/rise detectors) only steps on fresh samples and divides by the real interval the sample covers. Getting this wrong produced a fake 10 Hz wobble that used to trip the dip detector at random.

- A bad read is replaced with the last good one rather than a zero, so a single dropped sample can never inject a garbage spike into the PID.

## 4. Flow: from a tiny pressure difference to a usable number

flow_sensor.c

The raw differential pressure is genuinely noisy. A centrifugal blower makes 1–30 Hz pressure pulses, the restriction itself makes broadband turbulence hiss, and the housing vibrates into the sensor ports. So the raw signal is cleaned up in stages:

- Low-pass filter, 10 Hz by default. Fast enough to keep the shape of a breath, slow enough to kill the hiss.

- Zero offset removal. The sensor has a small offset that drifts. The code watches for genuinely quiet periods and re-learns the zero; you can also force it from USB or at therapy stop.

- Lookup table converts the cleaned pressure difference into litres per minute (slm).

- Spike watch. Sudden jumps are latched with a full snapshot of what else moved at the same moment, so you can tell afterwards whether a needle came from the air path or from the arithmetic.

A step-rejection rule freezes the filter on sudden jumps, but with an important exception: a real disconnect also produces a sudden jump, so the filter is allowed to catch up at about 600 slm/s. A rejection rule that hid mask-off would be worse than the noise it removed.

There are two flow numbers and they are not interchangeable:

- Flow_GetSlm() — total flow. Everything leaving the blower. Raw, no leak model involved. This is what mask-off detection uses, deliberately: it cannot be corrupted by a leak estimate that is still settling.

- Flow_GetPatientSlm() — patient flow. Total flow minus the vent flow minus the estimated leak. This is what the breath detector, EPR and the screen use.

If you are reading a plot and wondering why flow "decayed to zero" after a disconnect: that will be patient flow, and what you are watching is the leak estimator learning that the extra flow is leak, not breathing. Total flow did not change.

## 5. Tube compensation: the pressure we control is an estimate

tube_comp.c

The sensor is in the device; the patient is at the far end of a two-metre hose. When air flows, the hose eats some pressure:

hose loss  =  R_lam × |Q|  +  R_turb × Q × |Q|
mask P     =  sensor P  −  sign(Q) × hose loss

In words: a bit of loss proportional to flow (laminar), plus more loss proportional to flow squared (turbulent). When air is flowing towards the patient the mask sees less than the sensor; when the patient breathes out, the mask sees more.

The PID closes the loop on this estimated mask pressure, not on the raw sensor. That means the thing being held constant is what the patient experiences, which is the point of the whole device. It also means the controller is trusting a model. If R_lam/R_turb are wrong for your hose, the loop will hold the wrong pressure, confidently. Typical values for a 15 mm slimline hose are around R_lam = 0.002, R_turb = 0.00018; the header has a calibration procedure that needs only the device and a plug.

Turning tube compensation off makes the estimate equal the raw sensor reading, and the loop then regulates device-side pressure.

## 6. Leak estimation

leak_estimator.c, leak_estimator_vivo.c

Every mask has holes in it on purpose — the vent — so the patient does not rebreathe their own CO₂. That intentional flow is not a fault. What matters clinically is unintentional leak: a bad seal, an open mouth, a hose off.

The book equation is:

Q_total = Q_patient + Q_vent + Q_leak

- Q_patient averages to roughly zero over several breaths (what goes in comes out), so a slow filter removes it.

- Q_vent follows an orifice law, K × sqrt(mask pressure), with K around 8.7 slm per √cmH₂O by default for an adult mask.

- Whatever is left over is the unintentional leak.

Two backends are implemented and selectable at runtime:

- Simple: exactly the equation above, with a 10 second filter.

- Vivo-style (the default): a bench-measured lookup table of expected leak against pressure, multiplied by an adaptive MeanRatio that is corrected once per breath from the imbalance between inhaled and exhaled volume. It settles faster and copes better with masks that do not match the orifice law.

Above about 24 slm of unintentional leak the device calls it "high leak", which matches the usual industry convention. High leak also blocks APAP pressure increases, because the algorithms that decide to raise pressure cannot trust their inputs when the circuit is leaking badly.

One subtlety worth knowing: during a disconnect, the leak estimator would happily learn that 60 slm of escaping air is "normal leak". Mask-off therefore freezes leak adaptation while it is active (Leak_SetDisconnectFreeze).

## 7. Pressure control — the core loop

blower_ctrl.c, BlowerCtrl_Step(), 200 times a second.

This is the heart of the device, so this section is the long one.

### 7.1 The job

Given a target pressure, choose a motor RPM. Do it fast enough that a breath does not visibly move the pressure, and smoothly enough that the patient does not feel the controller working.

The output is a single number — RPM — posted to CM4 every tick.

### 7.2 Step one: slew the setpoint

The user-facing setpoint can jump instantly (ramp, EPR, a GUI change). The setpoint that actually feeds the loop is rate-limited to 18 cmH₂O per second in both directions.

Why: a step demand asks the loop to do something physically impossible, and the integrator winds up trying. Slewing the target means the error stays realistic and the loop stays in its linear region.

### 7.3 Step two: pick the gains (gain scheduling)

A blower is not a linear device. Near 4 cmH₂O, a small RPM change moves the pressure a lot. Near 20 cmH₂O, the same RPM change barely moves it. One set of gains cannot work across the whole range: tuned for the top it is sluggish at the bottom, tuned for the bottom it oscillates at the top.

So there are three sets, chosen by the setpoint:

| Zone | Range (cmH₂O) | Kp | Ki | Kd |
| --- | --- | --- | --- | --- |
| Low | 4 – 8 | 1800 | 4500 | 2.0 |
| Mid | 8 – 15 | 1200 | 4200 | 1.5 |
| High | 15 – 25 | 1000 | 3900 | 0.8 |

These came from an open-loop step-response fit on the bench (a first-order plus dead-time model: gain 0.000934 cmH₂O/RPM, time constant 68 ms, dead time 9 ms) run through the SIMC tuning rules, then adjusted by hand.

Scheduling is on the setpoint, not the measurement, and there is a 0.5 cmH₂O hysteresis band. If it scheduled on the measurement, a pressure sitting right on a zone boundary would flap between two gain sets several times a second.

### 7.4 Step three: the measurement

The feedback value is the tube-compensated mask estimate, lightly smoothed (roughly half old value, half new, each tick).

There is one refinement: during an inhale, if the pressure is sagging below target by more than a threshold, the loop blends towards the unfiltered measurement in proportion to how big the sag is. Filtering costs time, and time is exactly what you do not have when a breath is pulling the pressure down. The blend is only active during inhalation sags, so it does not add noise anywhere else.

### 7.5 Step four: P, I and D

P — proportional. Kp × error. The workhorse. It has an asymmetric twist:

- If the patient is inhaling and pressure is below target by more than 0.08 cmH₂O, P is multiplied by 3.5.

- If pressure is above target by more than 0.12 cmH₂O, P is multiplied by 6.0.

Both directions are hurry-up terms. Overshoot is punished harder than undershoot because an over-pressure lingers — the air has to bleed out through the vent — whereas an under-pressure is corrected by the blower itself.

I — integral. Removes steady-state error: the slow, patient accumulation that finds the exact RPM needed to sit on target. It has three protections, covered in 7.7.

It also has a deadband (1.2 cmH₂O by default): outside that band the integrator is frozen. The integrator's job is fine trim near the target, not powering through a large transient — that is P's and feed-forward's job.

And it integrates a filtered error rather than the raw one, with a filter that starts fast and slows down over the first minute of therapy. A slow filter from the very first tick would carry stale error from the previous session into this one.

D — derivative. −Kd × rate of change of the measurement. Note: of the measurement, not of the error. If it used the error, every setpoint change would produce a violent kick. Using the measurement, a setpoint change is invisible to D. The derivative is filtered (alpha 0.25) because raw differentiation of a 200 Hz signal is mostly noise.

### 7.6 Step five: feed-forward — the part that does the real work

Here is the key idea. A PID is reactive: it can only respond to an error that has already happened. By the time the pressure has sagged enough for P to notice, the patient has already felt it.

Feed-forward is predictive: it works out the RPM the blower should need, from physics, before any error exists. The PID is then left with only the small leftovers. This is why the gains above look modest — they are not carrying the load.

There are several feed-forward terms. Which are active depends on the profile.

The steady-state model (always the backbone). The relationship between fan speed and pressure is close to a square law, so inverting it gives:

RPM = 6141 × sqrt(pressure) + 2000

The 6141 came from a bench sweep of RPM against pressure — the scatter across the whole 4–25 cmH₂O range was under 0.5%, which is unusually good for a fan.

The +2000 offset matters and is worth understanding. The sweep was done with the outlet open to atmosphere, where the blower moves almost no air. In a real circuit the blower is always pushing the vent leak, so it needs more RPM than the open-air curve predicts — on the bench, about 2000 RPM more at 8 cmH₂O. That gap must be closed with an additive offset, not by raising the coefficient, because the coefficient is multiplicative and would also stretch every breath excursion and every RPM floor derived from it. An earlier attempt to raise 6141 to 6820 made the expiratory pressure hump noticeably worse.

Physics feed-forward (default ON). The steady-state model above assumes no flow. Extend it with the two things that change when air actually moves:

equivalent pressure = setpoint + hose loss(Q) + c_fan × Q²
RPM = 6141 × sqrt(equivalent pressure) + 2000

- hose loss reuses the same hose coefficients as tube compensation.

- c_fan × Q² is fan-curve droop: a fan produces less pressure when it is moving more air. Default 0.003, clamped at 0.005.

c_fan is the single biggest lever on the inspiratory dip. At the old value of 0.00012 the feed-forward anticipated only about 570 RPM of the ~2500 RPM a breath needs, so the P term supplied the rest — reactively, after the pressure had already dropped. To measure it properly: run the blower open-loop at a fixed RPM, sweep the flow, and plot pressure against flow squared. The slope is c_fan.

The flow that feeds this model is itself filtered, with different cutoffs for inhale (12 Hz, needs to be quick) and exhale (10 Hz).

Onset lead. Waiting for the flow to be large means arriving late. This term watches the rate of change of flow and adds an extra boost the moment a breath starts, before the flow peaks — up to 2500 RPM. The rate signal has to be filtered (30 ms) or its own noise produces a permanent phantom boost of a couple of hundred RPM.

Dip feed-forward (ON). If the sensor pressure is falling faster than 2 cmH₂O/s, add RPM in proportion, up to 1500. This is a backstop that works directly from pressure, so it still fires when the flow-based terms are confused about which way the patient is breathing.

The threshold has to sit just above the noise floor, not above the breath. The original 10 cmH₂O/s never fired once, because a 25 slm bench breath only produces about 1.6 cmH₂O/s. If the expiratory baseline looks hairy, this threshold has fallen into the noise — raise it to about 4.

Rise feed-forward (ON). The mirror image: when pressure is above the setpoint and still climbing faster than the threshold, subtract RPM. Nothing used to brake a rising pressure, and the physics model has nothing useful to say about it, because the real disturbance is the patient emptying their lungs into the circuit's compliance, which can only bleed out through the vent. So this term brakes on the measured pressure rate instead of on any model. It is deliberately not gated on breath phase, because the worst of the rise happens exactly at flow reversal, where the phase detector is least certain.

Expiratory flow feed-forward (ON). Fitted from real captures: the leftover over-pressure during exhalation is proportional to expiratory flow, at about 23 RPM per slm. But cross-correlation showed the pressure excess happens 200 ms before the flow — the flow is a correlate, not the cause. So this term uses a phase lead (Q + 0.2 × dQ/dt) rather than instantaneous flow, or it would arrive 200 ms late. It cannot remove the residual completely; flow only explains about half the variance.

Profiles. Two presets rearrange which terms run:

- CPAP Plateau: square-law feed-forward only, everything else off, softer gains, heavier filtering. Gives the flattest pressure platform.

- CPAP Hybrid: physics feed-forward during inhalation, plateau-style hold during exhalation.

They are mutually exclusive, and BlowerCtrl_EnforceFfMutualExclusion() stops incompatible combinations from being set by accident.

### 7.7 Step six: anti-windup, in three layers

Integral windup is the classic way a controller like this misbehaves: the output is already at maximum, the error persists, so the integrator keeps growing, and when the error finally clears the loop overshoots enormously because the integrator has to unwind first.

There are three independent defences:

- Conditional integration. Compute what the output would be. If it is saturated and the integrator step would push it further into saturation, skip the step.

- Headroom clamp. Limit the integrator to the RPM range actually left over after the feed-forward, P and D terms have taken their share.

- Hard magnitude clamp. An absolute ceiling on the integrator, scaled by the setpoint.

The integrator is also frozen while the slewed setpoint is still catching up to the user setpoint, since the error during that climb is expected, not a fault of the loop.

### 7.8 Step seven: clamp, filter, post

The terms are summed, clamped to the allowed RPM range (0 to 30000, where the ceiling is the safety limit), and passed through an output filter — though that filter is disabled by default (alpha 1.0), because the measurement filter and the term slew limits already do the smoothing.

Then the number goes to CM4 via BlowerIpc_CM7_SetSpeedRpmSilent(). "Silent" means no debug print; at 200 Hz the logging would flood the USB link. The ramp time is deliberately zero: a high-rate control loop wants each new target applied immediately, and any smoothing should happen inside the control law where it can be reasoned about, not inside the motor controller.

### 7.9 The whole loop in one line

RPM = feed-forward(setpoint, flow)        <- predicts most of the answer
    + Kp × error × (hurry-up factor)      <- reacts to what is left
    + integral(filtered error)            <- removes the last steady bias
    − Kd × d(measurement)/dt              <- damps
    + dip/rise/expiratory-flow terms      <- catches the breath transients
  then clamped and posted to CM4

### 7.10 WM6850 reference mode

Setting g_blower_ctrl_wm6850_mode = 1 switches the loop to imitate a reference implementation: 100 Hz, no feed-forward, no setpoint slew, no gain zones, no tube compensation, integrator starting at zero. It exists purely for A/B comparison against known behaviour, and it bypasses most of this document.

## 8. EPR — expiratory pressure relief

epr_ctrl.c

Breathing out against pressure is the main comfort complaint with CPAP. EPR drops the pressure by a fixed amount (1, 2 or 3 cmH₂O) while the patient exhales.

The convention here follows AirSense: the baseline pressure is CPAP − relief, and full CPAP is delivered only during confirmed inspiration. That is worth repeating because it surprises people on the bench: with EPR on and nobody breathing, the device sits at CPAP − relief, not at CPAP. That is correct behaviour, not a bug.

Breath phase comes from a small state machine on patient flow:

IDLE -> (sustained inspiratory flow) -> INSPIRATION
     -> (sustained expiratory flow)  -> EXPIRATION

Both transitions need the flow to persist, and inspiration must reach a minimum peak, so a cough or a leak wobble does not flip the pressure. Full CPAP is only applied in the confirmed INSPIRATION state — the "pending" state proved far too twitchy and caused visible pressure flapping.

Two types: full time (EPR always active) and ramp only (EPR during the ramp, then plain CPAP).

At therapy start there is a pressurisation hold: EPR is pinned at the expiratory pressure until the circuit has actually reached it and stayed there for a second, plus a 5 second startup grace where breath detection is ignored. This stops spin-up flow transients from being mistaken for breaths. Note that this hold is deliberately not applied when resuming after a mask reconnect — doing so held the resumed pressure a relief level low for five seconds and made the recovery look stalled.

## 9. The therapy state machine

therapy.c

IDLE -> RAMP -> RUNNING -> STOPPING -> IDLE
            \      |
             \     v
              -> FAULT (needs stop + start to clear)

- RAMP (optional, 0–45 minutes): start at 4 cmH₂O and climb linearly to the session target, so the patient can fall asleep at a low pressure. The climb is so slow (~0.01 cmH₂O/s) that the PID's own slew limiter never notices it.

- RUNNING: hold the target. In CPAP that is a fixed number; in Autoset the APAP algorithm moves it.

- STOPPING: a two-second spin-down.

- FAULT: blower disabled and stopped. There is no automatic recovery.

The session target is snapshotted at start so that GUI changes cannot yank a running ramp around, but the running target is re-read each tick so a deliberate change during therapy takes effect.

The state machine also collects session statistics — a pressure histogram producing the median, 90th and 95th percentiles, run time, mask-off and reconnect counts.

## 10. Mask-off detection

therapy.c. This is under active development; the behaviour below is current as of this document.

The problem. If the patient takes the mask off mid-session, the circuit is open. The pressure collapses, so the controller does the only thing it knows and drives the blower to maximum — loud, pointless, and it makes reconnecting unpleasant. So the device should notice, park the blower at a low speed, and resume properly when the mask goes back on.

Parking. On detection: disable the PID, ramp the motor down to 6000 RPM (soft, to avoid the deceleration pumping the DC bus into over-voltage), freeze leak adaptation, and hold. 6000 RPM is chosen so that there is still enough airflow to sense a reconnection.

Detecting the disconnect. Three independent paths, any of which can fire:

- Flow path. Total flow above 45 slm for 600 ms (300 ms if it is very high), together with collapsed pressure. Fast and obvious for a hose disconnect.

- Leak-ratio path. The Vivo estimator's 10-second mean ratio above a disconnect limit for 8 seconds. Slower, catches gradual seal failures.

- Pressure-only path. Commanded RPM at or above 12000 while both the sensor and the mask estimate sit below min(0.55 × setpoint, 4.0) cmH₂O, for 2 seconds. This reads no flow at all, so it still works when the flow signal is unusable or when the event is too small for the 45 slm threshold.

All three require the pressure to have actually collapsed, so a strong inhalation cannot trigger a mask-off.

Why the third path exists. A mask lifted off the face is a much smaller event than a hose disconnect: the vent orifice limits the flow. Measured on this bench at a setpoint of 8, the circuit still holds about 2.5 cmH₂O and passes only about 20 slm with the blower pinned at 30000 RPM. That is below the 45 slm flow threshold and above a fixed low-pressure floor, so neither of the first two paths fires. Hence a threshold that scales with the pressure being asked for.

Detecting the reconnect. At 6000 RPM, an open hose reads near zero pressure and a sealed mask tips up to roughly 0.7 cmH₂O. So the device learns an open-circuit reference once the deceleration has settled, then watches for the pressure to rise above it. There are two fallbacks:

- Sustained flow collapse, for the case where the mask went back on before the deceleration finished — the reference is then the sealed pressure itself, and no rise can ever appear above it.

- A 60 second watchdog that re-probes at therapy pressure, so a mask-off can never latch for the rest of the session no matter what the sensors do.

Resuming. Re-seed the controller at the steady-state RPM for the target, floor the RPM for 5 seconds so it cannot collapse, suppress patient-flow feed-forward for 3 seconds (the leak estimator is re-seeding and briefly mistakes hose-filling flow for the patient), reset the breath state, and give the safety supervisor a 12 second grace while the pressure climbs back.

Known limitation. Detection is reliable from about 5 cmH₂O upwards and gets better as the setpoint rises. At the very bottom of the range it degrades, because the open-circuit pressure is set by the blower's speed limit and the circuit — not by the setpoint — so at a setpoint of 4 the difference between "mask on" and "mask off" is only about 1.5 cmH₂O. A pressure-based test cannot reliably separate those. The robust fix is to compare commanded RPM against the RPM the square-law model says should deliver this pressure, which is independent of the pressure scale; that is not implemented yet.

## 11. APAP / Autoset

apap_ctrl.c

In Autoset mode the device chooses its own pressure between a minimum and a maximum. This is a functional prototype, not a clone of anyone's algorithm, and the thresholds are heuristics awaiting clinical validation.

The loop:

- Start at the minimum pressure.

- Monitor while breathing continues. Look for flow limitation (a flattened or peak-dropped inspiratory shape), hypopnea, and snore. Each raises the pressure by a small or moderate step depending on severity.

- No breath for about 4 seconds → start the forced-oscillation test (below). Score the event only if it lasts around 10 seconds, so that a small obstructed gasp does not abort the measurement.

- Classify: if the pressure oscillation gets through but the flow oscillation does not, the airway is closed — an obstructive apnoea, so raise the pressure in duration-based steps. If flow is transmitted, the airway is open and the patient simply is not breathing — a central apnoea, and raising pressure would not help, so it does not.

- High leak blocks every pressure rise. If the circuit is leaking, the event detectors cannot be trusted.

- After a quiet, stable period, drift the pressure slowly back down.

A "Soft" response setting makes all the rises gentler.

### Forced oscillation (FOT)

blower_osc.c

To tell an open airway from a blocked one when the patient is not breathing, the device adds a small sine wave (about 4 Hz, around 1 cmH₂O peak) to the pressure setpoint. The PID reproduces it, so the pressure in the circuit gently oscillates.

If the airway is open, that oscillation produces a matching oscillation in the flow. If it is blocked, the pressure wiggles but the flow does not. The ratio of the two is the airway impedance.

Implementation details that matter: the phase is kept as an accumulator in [0, 1) so it cannot drift after hours of running; the sine always starts at a zero crossing so switching it on does not step the pressure; amplitude and frequency are clamped so a typo in a watch window cannot harm a patient.

The FOT signal must be scored from the raw flow, not the published patient flow — the breath-detection filters on the patient-flow path remove the 4 Hz carrier entirely.

## 12. The safety supervisor

safety_supervisor.c, its own task, highest priority.

It watches five things and latches a fault on any of them. A fault disables the controller, stops the motor, and requires an explicit stop-then-start.

| Fault | Condition | Meaning |
| --- | --- | --- |
| Low pressure at high RPM | Sensor below 1.5 cmH₂O while commanding over 12000 RPM, for 4 s | Hose off, sensor dead, or huge leak |
| Over pressure | More than 5 cmH₂O above target for 3 s | Sensor stuck high or blockage |
| Sensor out of range | Outside −10…+50 cmH₂O for 2 s | Electrical fault |
| RPM saturated | Pinned at the 30000 RPM limit without reaching target, for 6 s | Controller cannot deliver — leak too big or blower undersized |
| CM4 heartbeat | Mailbox status counter stopped for 100 ms | Motor core dead |

Three exemptions are built in, and they exist for real reasons:

- A start grace (10 s) and a reconnect grace (12 s), because the pressure is legitimately low while the circuit fills.

- While mask-off is active, because parking at 6000 RPM with near-zero pressure is intended behaviour, not a fault.

- While a mask-off detector is arming, because the fault timers can start counting earlier than the detectors (a large leak pins the RPM well before a full disconnect drops the pressure to the floor), and without this the fault would stop the motor while the detector was still confirming.

That last exemption is narrow on purpose. A genuine "cannot reach target" leak never puts the pressure low enough to arm mask-off, so it still faults exactly as intended.

## 13. The motor side

CM4/Core/Src/blower_ipc_cm4.c plus the ST motor control library.

CM4 runs field-oriented control, which is the standard way to drive a brushless motor smoothly: it measures the phase currents, works out where the rotor is, and computes the voltages that produce the torque wanted with the least ripple. The details live in ST's generated library; this firmware mostly configures it and feeds it a speed target.

Two pieces of local logic matter:

Auto-arming. CM7 often posts a speed command without a preceding start, so a SET_SPEED arriving while the motor is idle or in a cleared fault will acknowledge the fault and start the motor itself. Without this, a lone speed command could leave the rotor coasting.

The regenerative brake. Slowing a spinning blower turns the motor into a generator and pushes energy back into the DC bus. The over-voltage trip is at 24.9 V, which is not far above the normal rail, so a fast deceleration can trip it and kill the PWM — leaving the rotor coasting exactly when control is wanted. A brake resistor on a PWM output dumps that energy: it engages above 22 V, scales with how far over the threshold the bus is, and releases at 21 V. Large planned decelerations (like the mask-off park) pre-engage it.

## 14. Humidifier

humidifier.c, on the slow ClimateTask.

A 25 W coil heats about 100 mL of water; an SHT4x measures humidity and air temperature above the bowl. Full power reaches roughly 50 °C in 10–15 minutes, so the thermal time constant is minutes with significant dead time.

The loop is therefore deliberately slow and unambitious: a PI controller at 1 Hz, conditional-integration anti-windup, a slew limit on the PWM output, hard duty clamps, an over-temperature trip, and a "heater stuck on" watchdog that faults if the PWM sits near maximum for a long time without the humidity rising.

Disabled at boot.

## 15. Tuning and observability

Nearly every constant mentioned here is a volatile global, which means:

- You can watch and change it live in the Keil watch window or STM Studio.

- You can change many of them over USB with therapy set ..., blower ... and similar commands (see telemetry_stream.c for the list).

- The names follow a pattern: g_<module>_<thing>, for example g_blower_ctrl_kp_low, g_therapy_mask_off_enter_slm.

Telemetry. A 50 Hz stream over USB CDC carries sensor pressure, mask pressure, patient flow, commanded RPM and actual RPM. pc_tools/cpap_telemetry_plot.py plots it live with a status panel. A 100 Hz stream goes to an ESP32 over UART for MQTT. Sessions can also be recorded to SD card in EDF format (see SESSION_EDF_LOGGING.md).

Suggested PID bring-up order, if you ever need to retune from scratch:

- Set Ki and Kd to zero. Raise Kp until you see mild oscillation, then back off about 30%.

- Raise Ki until steady-state error clears in 200–500 ms.

- Add a small Kd (5–30) only if there is overshoot you cannot remove by lowering Kp.

Do it with the feed-forward configured the way you intend to ship it. The gains assume feed-forward is carrying the steady-state load; with it off you will need two to three times more Kp and Ki.

## 16. Gotchas worth knowing before you debug something

- The two filter alphas mean opposite things. g_blower_ctrl_meas_filter_alpha weights the old value (0.45 = light smoothing). g_blower_ctrl_out_filter_alpha weights the new value (1.0 = no filtering at all). This has caught people out.

- BlowerCtrl_GetOutputRpm() freezes when the PID is disabled. During mask-off or a fault, BlowerCtrl_Step() returns before it updates the output registers, so that value keeps whatever it last commanded. Mask-off now publishes the parked RPM explicitly to keep it honest, but anything else that disables the loop will leave a stale number on the screen and in the logs.

- Patient flow is a derived number. If it does something surprising, check total flow first. Most "flow did something impossible" moments are the leak estimator adapting.

- The mask pressure is an estimate, not a measurement. Nothing in the device actually measures pressure at the mask.

- Total flow, not excess flow, drives mask-off. This is deliberate. Excess flow depends on the leak and vent estimates, which are exactly the things that go wrong during a disconnect.

- Setpoint zones schedule on the setpoint, not the measurement, so during a big transient the gains in use are the ones for where you are going, not where you are.

- A dropped sensor sample repeats the previous value. Anything that differentiates must check g_ps_sample_fresh, or it will see a zero derivative that is not real.

## 17. Where to look in the code

| Topic | File |
| --- | --- |
| Pressure PID and feed-forward | CM7/Core/Src/blower_ctrl.c |
| Mask pressure estimate | CM7/Core/Src/tube_comp.c |
| Flow conditioning | CM7/Core/Src/flow_sensor.c |
| Leak estimation | CM7/Core/Src/leak_estimator.c, leak_estimator_vivo.c |
| Therapy states, mask-off | CM7/Core/Src/therapy.c |
| EPR | CM7/Core/Src/epr_ctrl.c |
| APAP / Autoset | CM7/Core/Src/apap_ctrl.c |
| Forced oscillation | CM7/Core/Src/blower_osc.c |
| Safety | CM7/Core/Src/safety_supervisor.c |
| Cross-core mailbox | Common/Inc/blower_ipc.h, CM7/.../blower_ipc_cm7.c, CM4/.../blower_ipc_cm4.c |
| Humidifier | CM7/Core/Src/humidifier.c |
| USB telemetry and commands | CM7/Core/Src/telemetry_stream.c |

The header files carry the design rationale — in this codebase they are worth reading before the implementation.